Empyrean Use and Procedures

XRD manager: David Tavakoli

david.tavakoli@mse.gatech.edu

Radiation Safety Class:  Submit your FORM B in Office of Radiological Safety. You can use www.ors.gatech.edu to take the class or refresher. Or go to:

ehs1.fac.gatech.edu/ehsaweb/ehsawebisapi.d11/?protocol=GTRAD

Please sign the logbook (Time In, Time Out, and Instrument Status).

# Radiation : The Empyrean in its default configuration emits  alpha (or Kα) and beta radiation, unlike the Alpha-1 which gets rid of secondary radiation and only has monochromated radiation (means only one single wavelength).

It Does the following analysis:

-Nanostructure analysis, line profile, particle size (strain measurement, rocking curve analysis, residual stress, and Small Angle X-ray (SAXS) measurements. This guide will only cover line profile analysis.

Instrument Basics.

The Empyrean with a reflection-transmission spinner in place is a basic theta-two-theta diffractometer (source rotates θ angle and detector rotates 2θ angle and the angle between the sample and the detector will remain as θ) with the benefit of an incident beam mirror, and a fast linear detector (Pixel Medipix 3D).  It can generate very high resolution diffraction patterns in a minimum amount of time, It’s a fast data collector using RTMS (Real Time Multiple Strip) as a detection technology.  The sample stays fixed and the x-ray ((line source on this instrument) source and detector rotate to collect the diffraction pattern.

The Empyrean has an incident beam mirror as a prefix module, a reflection-transmission spinner for a sample stage, and diffracted beam optics. Each of these will need to be configured to suit your particular measurement.

- Inside the XRD- Instrument (Mechanically inserted optics)

Open XRD door

-	Incident Beam Optics Setup

1. Figure out what your starting angle will be. Then choose the slits accordingly using the excel spreadsheet, BBHD Setup Guide), as it will assist you in determining what slits to use.  All optics except for the masks are located in the bin on the back right of the instrument. The masks are located in the wooden holder beneath the BBHD.

2. The Divergent Slit is always ¼ the size of the Anti-Scatter Slit.

3. Soller slit. It is an optic with several parallel thin brass sheets. Choose either the 0.04 (recommended) or 0.02 radian soller slit depending on sample size. DO NOT DROP IT!!! Grab the handle then insert it. It relates to the sharpness of the peak and the resolution.

4. Mask (Filter). Choose the mask to fix the beam size according to your sample diameter.

-	Diffracted Beam Optics Setup. (Receiving Optics)

5. The diffracted beam side is easy. It has one slit ... unfortunately it's also called an Anti-scatter slit.  For the most part, the 7.5 or 8 degree anti-scatter slit will be in place and will not need to be adjusted by the user.

- Sample Preparation

You will either load a powder mount, a wafer mount or a low-background mount. That's it.  It is vital to ensure your sample is positioned exactly at the top of the powder and wafer mounts. People using the low-background mounts should only have a monolayer of powder.  These samples are not designed for thick (greater than 100micron) samples.

6. Load the sample in the sample spinner stage and use the software to load the sample in position.

XRD SETUP

Classical powder diffractometer setup in Bragg-Brentano geometry

This geometry is the most commonly used geometry for phase analysis for reflection measurements.

The focal spot of the tube, the sample surface and the detector slit are positioned on the so-called focusing circle.

Effect of sample displacement with focusing geometry

III- Computer set up

-	 X-pert  Datacollector  program  should be  ON

7. The program should be already logged onto and connected to the instrument.

8. Verify the Incident Beam Settings by clicking on Incident Beam Optics menu then double click on the first row (incident beam path)- The Incident Beam Optics window will open and recheck all the optics. Make sure the optics are what you want. Click on: Anti-scatter slit, mask, soller,… and choose the desired values.

On the same window:

Click on Divergence slit. You can either choose Fix or Automatic option. If you choose Automatic option, Do not change any number. If you choose Fixed, then the value is half of your Anti-scatter slit.

Other optics should be on:

prefix module : BBHD

Mirror : BBHD

Click Apply and then OK

9. Do the same procedure to verify the Diffracted Beam Optics, by clicking on Diffracted Beam Optics and double click on Diffracted Beam Path1. A window will open and recheck all the optics.

Other optics should be on:

PreFix module : X'Celerator (wide angle).

Detector : Actual

Monochromator, Receiving slit, and Collimator : None

Apply then OK

10. Go to the Instrument settings menu, double click on current then Increase the current to 40mA and keep the voltage at 45KV                                                                                      - Check Generator On, shutter is still closed.

-	Setting up a Measurement.

11. Go to File>New Program. Use program type "Absolute Scan" (if you already have a program, you may use it).

12. Choose the angle ranges you desire. Example 20-80 degrees, Gonio scan, and continuous

13. Choose a step size and count time... these are rolled into a scan rate.  General settings are 0.016 degrees (step size) and 10 seconds (time per step, very fast). Don't worry if it changes the numbers on you to the closest value that the hardware is capable of depending on what you choose. Scan speed and Total time will be calculated by the program.

To see the actual time for the scan run, at the same window (Absolute scan window), go to setting>click on the option “Detector-Actual” and change it to PIXcel Medipix then OK.

14. Close window, Yes to Save the program.- Name it.

-	Making a Measurement

15. Go to Measure>Program. Open your program.

16. Give your file a name, and a directory.

17. Verify that the instrument is ready (status: 45 KV, 40 mA, the doors are closed, and your sample is raised up into position).

18. Hit "OK" to begin your measurement. You will hear a click. The shutter is now open (cps vs. 2 θ angle) and you see the scan plot. You should see a RED Stop SIGN on the top of your computer screen warning you to “Not open the X-ray door”. When the scan is DONE you will hear a click, and now you the data is saved and the run is completed.

IV-Return to stand-by when measurement is complete

Lower the power of the x-ray tube (Decrease the Current to 20 mA. Voltage remains at 45 kV) and OK.

Open Xray door and pick up your sample.

Scheduling on line system with GTLogin and Password on SUMS.

https://sums.gatech.edu/Dashboard.aspx

To convert the file:

You can use Data viewer to convert the file.						      	         Open up .xrdml format file from your folder

(ex. My computer > OS> c:\x’pert data > File name)

- File> convert   (from the top menu)

- Check mark the box (csv for excel file) then click convert

- Close the file then go back to your folder again and you should see the excel formatted file.

Please clean up the sample area and inside the XRD after you are finished. And please write down the Time OUT and instrument status in the logbook.